Steven A. Jones

BIEN 557, Medical Ultrasonics

August 25, 2004


Exercises:
1. Estimate the Doppler shift caused by an ambulance as it passes an observer at 30 mph.  Assume that the ambulance emits a sound with frequency of 5 kHz. Let the speed of sound in air be 33,000 cm/s.  What is the appropriate Doppler angle for this problem?  Why?

Solution:

· From the Doppler equation, 
[image: image31.jpg]


, one obtains 
[image: image2.wmf](

)

Hz

cm

Hz

in

cm

ft

in

mi

ft

hr

hr

mi

4

.

406

0

cos

sec

/

000

,

33

000

,

5

54

.

2

12

5280

sec

60

min

1

min

60

1

30

2

=

÷

÷

ø

ö

ç

ç

è

æ

´

÷

÷

ø

ö

ç

ç

è

æ

´

´

´

´

´

´

=

o

d

f


The Doppler angle to use is 0 degrees because one is looking at the difference in between the frequency observed as the ambulance is approaching at 0 degrees and then receding (at 180 degrees).  To be precise, we the total shift will be:


[image: image3.wmf](

)

(

)

180

cos

0

cos

0

0

c

vf

c

vf

f

d

-

=

, where the first term is the shift in the signal as the ambulance is approaching and the second term is the shift when the ambulance is receding.  It is just fortunate that 
[image: image4.wmf](

)

1

180

cos

-

=

, so that 
[image: image5.wmf](

)

(

)

(

)

0

cos

2

180

cos

0

cos

0

0

c

vf

c

vf

c

vf

f

o

d

=

-

=

.
2. Expand Equation 2 through a Taylor series up to second order in velocity.  By examining the second order term, estimate the error in the Doppler frequency caused by the narrow band assumption used in Equation 3.  Assume:


[image: image6.wmf]cm/s

100

=

v



[image: image7.wmf]cm/s

000

,

150

=

c



[image: image8.wmf]MHz

10

0

=

f



[image: image9.wmf]o

30

2

1

=

=

q

q


Solution:

Equation 2 is:


[image: image10.wmf]÷

ø

ö

ç

è

æ

-

=

2

cos

1

/

q

c

v

f

f

t

r

.  Letting 
[image: image11.wmf](

)

2

cos

q

c

v

x

=

, and recalling that 
[image: image12.wmf](

)

(

)

(

)

(

)

K

+

¢

¢

+

¢

+

=

0

2

0

0

2

f

x

f

x

f

x

f

, one finds that the second order term is 
[image: image13.wmf](

)

t

t

f

f

c

v

2

2

2

000

,

150

100

2

1

cos

2

1

÷

ø

ö

ç

è

æ

=

÷

ø

ö

ç

è

æ

q

.  The Doppler shift, however is 
[image: image14.wmf]t

f

000

,

150

100

2

´

, so the second order term is 
[image: image15.wmf]00267

.

0

1500

/

4

=

 times the Doppler frequency.  Thus, by ignoring the second order term one makes an error of only about 3 parts in 1000, which is far smaller than the error caused by other factors, such as spectral broadening and coherent scattering.
3. Use standard trigonometric identities to show that when 
[image: image16.wmf][

]

(

)

f

w

w

+

+

t

A

d

r

0

cos

 is multiplied by 
[image: image17.wmf](

)

t

A

t

0

cos

w

, the result can be written in the form of Equation 5:


[image: image18.wmf](

)

(

)

[

]

(

)

f

w

f

w

+

+

+

+

=

t

w

A

A

t

A

A

t

s

d

r

t

d

r

t

0

2

cos

2

1

cos

2

1


4. Verify that the phase 
[image: image19.wmf]f

 in Equation 5 is the same as the phase 
[image: image20.wmf]f

 in 
[image: image21.wmf][

]

(

)

f

w

w

+

+

t

A

d

r

0

cos

.

Since there were no changes of variables in the derivation of the previous problem, the phases are the same.

5. Use MatLab to generate and plot the signals on the left hand side of Figure 3.  Specifically, generate 
[image: image22.wmf](

)

t

A

t

0

cos

w

 and  
[image: image23.wmf][

]

(

)

f

w

w

+

+

t

A

d

r

0

cos

, and then multiply them together in the time domain.  Then use a low pass filter to obtain the downmixed Doppler signal.  Use 
[image: image24.wmf]kHz

1

0

=

w

, 
[image: image25.wmf]Hz

d

100

=

w

, 
[image: image26.wmf]1

=

=

r

t

A

A

, and 
[image: image27.wmf]0

=

f

.
This was done in class.

6. Compare the calculated result in 5 to the result of calculating the downmixed signal directly from Equation 5 and calculating the filtered downmixed signal directly from the first term on the left hand side of Equation 5.
The two should compare closely, assuming that a good filter was used.  However, there will likely be edge effects and possibly a phase shift caused by the low pass filtering.
7. Use the fast Fourier transform routine in MatLab to calculate and plot the power spectrum of the time-domain signals calculated in problem 5.  Are the spectra infinitely narrow, as in a delta function?  Why or why not?
The spectra will not be infinitely narrow because the signals are not infinite in time.

8. Repeat problem 7, but before taking the Fourier transform zero pad the signal to quadruple the number of points.  How would you describe the shape of the power spectrum mathematically?

The peaks should start to look like sinc functions.  It is as if you multiplied a longer signal by a “boxcar” shaped signal (1 for half the time and then zero for the remainder of the time).  Thus, the spectrum is a convolution of the sinc function (from the boxcar) with the delta function (from the sine wave).

9. If both the Fourier transform of a signal and the downmixing of a signal are linear processes, explain why the power spectrum of a Doppler signal from two simultaneous scatterers is not equal to the sum of the power spectra of the individual scatterers.
The Fourier transform is linear, but the power spectrum is the magnitude squared of the Fourier transform and is therefore not linear.

10.  Use MatLab to generate all of the signals on the left hand side of Figure 5.
11. Use MatLab and the FFT to generate all of the spectra on the right hand side of Figure 5.

12. Use MatLab to generate the three time-domain signals in Figure 6.

13. Use MatLab and the FFT to generate all of the spectra in Figure 6.

14. Show that if the transmitted signal pulse has a Gaussian shape to it and the receiver gate shape is also Gaussian, then the dependence of the amplitude of the signal from a particle within the sample volume on axial location is also Gaussian.
15. Use MatLab to calculate the first moment, mode, and max frequencies of the Doppler spectrum from a continuous wave downmixed signal (as calculated in problem 8).

16. Use MatLab to add random noise to a cosine wave (signal to noise ratio 20 db, i.e. the rms of the noise should be 1/10th of the rms of the cosine).  Plot the resulting signal.

17. Calculate the first moment, mode and max frequencies of the signal obtained in 16 above.  Compare these frequencies and explain differences, if any.

18. Create a zero crossing detector in MatLab.  Apply this detector to the signal generated in 16 above and compare the result to the frequencies obtained in 17.
19. For a typical 10 MHz Doppler device used in carotid diagnosis, with the pulse repetition frequency equal to 32 kHz, calculate the velocity at which aliasing occurs and the spacing between the multiple sample volumes.

20. Give mathematical formulations for each of the parameters 
[image: image28.wmf](

)

r

G

r

, 
[image: image29.wmf](

)

t

r

a

,

r

, and 
[image: image30.wmf](

)

t

r

v

,

r

r

 in Equation 9 for a typical Doppler ultrasound simulation.

21. Use MatLab to calculate and plot the beam pattern for the far field of a circular transducer element of a 10 MHz device for the following conditions:

a. transducer diameter = 0.5 mm

b. transducer diameter = 1 mm

c. transducer diameter = 2 mm

22. Derive mathematically the expressions for the spectra in panels a, b and c of Figure 9.
23.  Estimate the total scattered power from a single red blood cell.
24. Plot scattered power as a function of scattering angle for a single red blood cell (model the cell as a sphere in this case).

25. Estimate the depth of penetration of a 2 MHz, 5 MHz, 10 MHz and 20 MHz Doppler ultrasound signal.  Assume that depth of penetration is defined as the depth at which the sound power is 10% of the transmitted power.  Based on this calculation, does the following rule of thumb apply:

a. Transcraneal Doppler: 2 MHz

b. Cardiac Doppler: 5 MHz

c. Carotid Doppler: 10 MHz

d. intravascular Doppler: 20 MHz

26. Estimate the percentage of power that penetrates to 1 cm for the 2 MHz, 5 MHz, 10 MHz and 20 MHz ultrasound devices.

[image: image1.wmf](

)

q

cos

2

0

c

vf

f

d

=

_1157460940.unknown

_1157461413.unknown

_1157462243.unknown

_1157462322.unknown

_1157523810.unknown

_1157462371.unknown

_1157462283.unknown

_1157461754.unknown

_1157461904.unknown

_1157461623.unknown

_1157461057.unknown

_1157461347.unknown

_1157461018.unknown

_1154945568.unknown

_1154945636.unknown

_1154946489.unknown

_1157460384.unknown

_1154946468.unknown

_1154945593.unknown

_1141545395.unknown

_1154945493.unknown

_1141547028.unknown

_1140024802.unknown

_1140024912.unknown

_1140109351.unknown

_1140024852.unknown

_1139943293.unknown

